- هدف: - هدف از تدوین این دستورالعمل، تشریح روند اجرای کار با دستگاهFT-IR ATR با نام تجاریBruker ساخت کشور آلمان می باشد. - دامنه عملکرد: - این دستور العمل در بخش های مختلف آزمایشگاه کاربرد دارد. - مسئولیت : - مسئولیت اجرای این دستور العمل با کارشناسان آزمایشگاه مربوطه می باشد. - 4- روش کار: - از روشن بودن قسمت اصلی دستگاه مطمئن شوید. - بعد از روشن کردن لب تاپ دستگاه، بر روی برنامه OPUS کلیک کرده و جهت ورود پسورد "OPUS" را وارد کنید. - اجازه دهید تا نرم افزار برنامههای آماده سازی را به اتمام برساند. - به نسبت 1/20 از ماده مورد آنالیز/KBr را داخل هاون ریخته و کاملا پودر نمایید. - قبل از گرفتن طیف نمونه مورد نظر، مزاحمت محیطی را با اندازهگیری پس زمینه حذف کنید. - بعد از آماده سازی نمونه، به کمک نگاهدارنده دستگاه، قرص ساخته شده را داخل دستگاه گذاشته و سپس طیف نمونه قرص را بگیرید. - بعد از اتمام کار و خاموش کردن لب تاپ، دستگاه را در حالت روشن رها کنید. - در صورت نیاز به استفاده از ATR، با کارشناس آزمایشگاه به منظور آماده سازی دستگاه هماهنگ کنید.

1. Purpose: The purpose of this procedure is to describe the operating procedure for the FT-IR ATR instrument, Bruker brand, manufactured in Germany. 2. Scope: This procedure is applicable to the different sections of the laboratory. 3. Responsibility: The laboratory personnel responsible for the relevant section are responsible for implementing this procedure. 4. Procedure: - Ensure that the main unit of the instrument is switched on. - After switching on the instrument laptop, open the OPUS software and enter the OPUS password to log in. - Allow the software to complete its initialization and preparation procedures. - Mix the material to be analyzed with KBr at a ratio of 1:20 and thoroughly grind the mixture in a mortar until it is completely powdered. - Before acquiring the spectrum of the sample, eliminate environmental interference by measuring the background spectrum. - After preparing the sample, place the prepared pellet into the instrument using the sample holder, and then acquire the spectrum of the sample pellet. - After completing the analysis and shutting down the laptop, leave the instrument switched on. If ATR measurement is required, coordinate with the laboratory technician responsible for preparing the instrument.