- هدف: - هدف از تدوین این دستورالعمل، تشریح روند اجرای کار با دستگاه Nano Sizer (ZS) با نام تجاری ساخت کشور انگلیس می باشد. - دامنه عملکرد: - این دستور العمل در بخش های مختلف آزمایشگاه کاربرد دارد. - مسئولیت : - مسئولیت اجرای این دستور العمل با کارشناسان آزمایشگاه مربوطه می باشد. - 4- روش کار: - روشن کردن دستگاه و سپس کامپیوتر - اجرا کردن نرم افزار مربوطه (Zeta Sizer version 7.11) - بعد از باز کردن درب دستگاه، با توجه به نوع اندازه گیری (Zeta/Size) cell مربوطه را انتخاب کرده و داخل دستگاه قرار دهید . - وارد قسمت measure و سپس manual دستگاه شوید . - در قسمت measurement type نوع اندازه گیری ( Zeta potential/Size) را انتخاب کنید . - در قسمت material نوع ماده خود را انتخاب کنید . - در قسمت dispersant فاز پراکنده کننده نمونه خود را انتخاب کنید . - در قسمت temperature دما و مدت زمان مورد نظر برای به تعادل رسیدن دما برحسب ثانیه را وارد کنید . - سپس در قسمت cell نوع cell مربوطه را انتخاب کنید. - در قسمت measure تعداد تکرار مورد نظر برای اندازه گیری دستگاه وارد کنید. - کلید OK را فشار دهید . - فلش سبز رنگ start از قسمت بالای صفحه باز شده را فشار دهید . - نکته مهم: - از باز کردن درب دستگاه زمانی که دستگاه در حال خوانش میباشد خودداری کنید. - هنگام وارد کردن نمونه در cell مورد نظر از سرنگ یا سمپلر استفاده کنید تا از ایجاد حباب جلوگیری شود . - نمونه را در cell مورد نظر تا خط نشانه مورد نظر پر کنید - پس از هر بار استفاده cell مربوطه را دو مرتبه با اتانول و سه مرتبه با آب مقطر بشویید

1- Objective: The purpose of this procedure is to outline the operational process for the Nano Sizer (ZS) instrument (UK-manufactured). 2- Scope: This procedure applies to various sections of the laboratory. 3- Responsibility: The relevant laboratory specialists are responsible for implementing this procedure. 4- Procedure: - Turn on the instrument, followed by the computer. - Launch the associated software (Zeta Sizer version 7.11). - Open the instrument lid, select the appropriate cell based on the measurement type (Zeta/Size), and place it inside the instrument. - Navigate to the "Measure" section and then select "Manual." - Select the measurement type (Zeta potential/Size) under the "Measurement Type" section. - Select the material type under the "Material" section. - Select the sample's dispersing phase under the "Dispersant" section. - In the "Temperature" section, enter the desired temperature and the time (in seconds) required for the temperature to stabilize. - Then, select the appropriate cell type in the "Cell" section. - In the "Measure" section, enter the desired number of measurement repetitions. - Press the "OK" button. - Press the green "Start" arrow located at the top of the screen. - Important Notes: - - Do not open the instrument's lid while a reading is in progress. - - Use a syringe or sampler when introducing the sample into the cell to prevent the formation of air bubbles. - - Fill the cell with the sample up to the designated mark. - - After each use, rinse the cell twice with ethanol and three times with distilled water.