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	عنوان مدرك:  دستورالعمل اجرایی دستگاه اسپکتروفومتر

Title: Spectrometer Perkinelmer EquipmentWorking Instruction


	كدمدرك:
WI- RC- 010
	بازنگری:
----



	1- هدف: 
هدف از تدوین این دستورالعمل، تشریح روند اجرای کار با دستگاه UV-Spectrophotometry با نام تجاریPerkinelmer ساخت کشور آمریکا می باشد. 
2- دامنه عملکرد: 
این دستور العمل در بخش های مختلف آزمایشگاه کاربرد دارد. 
3- مسئولیت :
مسئولیت اجرای این دستور العمل با کارشناسان آزمایشگاه مربوطه می باشد. 
4- روش کار:
1- روشن کردن دستگاه و سپس کامپیوتر
2- قراردادن کووتهای (سل) پر شده از بلانک در محل مربوطه در نگهدارنده کووت در داخل دستگاه
3- Run کردن نرم افزار مربوطه (Win lab) و زدن دکمه Auto Zero و سپس زدن دکمه start (دوبار)
4- حال دستگاه آماده اندازهگیری نمونه است، لذا سل موجود در جایگاه sample را برداشته، از بلانک خالی کرده و نمونه را داخل آن میریزیم و بعد از قرار دادن در محل و بستن درب دستگاه مجدداً دکمه start را میزنیم  
5- بدین ترتیب میزان جذب نوری نمونه و طیف اسکن آن بدست میآید. 
نکته مهم:
· از باز کردن درب دستگاه زمانی که دستگاه در حال خوانش میباشد خودداری کنید.



	1- Objective:
The purpose of this procedure is to outline the operational steps for the PerkinElmer UV-Spectrophotometer (manufactured in the USA).
2- Scope:
[bookmark: _GoBack]This procedure applies to various sections of the laboratory.
3- Responsibility:
The relevant laboratory specialists are responsible for implementing this procedure.
4- Procedure:
1. Turn on the instrument, followed by the computer.
2. Place the cuvettes (cells) filled with the blank solution into the designated slots within the instrument's cuvette holder.
3. Launch the associated software (WinLab), click the "Auto Zero" button, and then press the "Start" button (twice).
4. The instrument is now ready for sample measurement; remove the cell from the sample position, empty the blank solution, fill the cell with the sample, place it back in the holder, close the instrument lid, and press the "Start" button again.
5. In this way, the sample's absorbance and its scan spectrum are obtained.
Important note:
- Do not open the instrument's lid while it is performing a reading.
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